Report on Australian Spinal Research Foundation Grant LG2004-03

Self testing of cortisol for well-being

27 May  2008

Aims of the study

The general aims of this study were stated in the application as follows;

- To provide a simple 5 minute accurate home testing system for tetrahydrocortisone-3-glucuronide (THEG) in timed urine samples.

- To provide a rapid screening test which can be used to monitor general personal well-being.

- To serve as an objective confirmatory test for the efficacy of chiropractic manipulation in asthma patients.

Associated questions to be answered were;

- Can the Ovarian Monitor be used as a home test for THEG?

- Can a quantitative immunochromatographic test for THEG be developed?

- Do the levels of THEG determined by the new tests correlate with urinary free cortisol levels in normal patients?

- Can elevated THEG levels be recognised by self testing or point-of-care testing  with the Monitor and/or immunochromatographic tests?

- Can home measurement be used to demonstrate that spinal manipulation reduces elevated cortisol/THEG levels to normal values in chronic disease situations.?

Clearly all of these aims and questions can only be tested or demonstrated once a test system for THEG is available. Equally clearly not all the goals were expected to be realised during the course of the current study given the levels of resources available. The first aim of the study therefore was the most important to deal with initially since without a viable testing system for THEG levels none of the other  aims or questions can be investigated.

Research Design

Under the Research design section of the application we stated that the goals and specific objectives were the following:

- synthesis of biomaterials for THEG assays

- validation of an homogeneous enzyme immunoassay (as used in the Ovarian Monitor) for urinary THEG levels against a method for urinary free cortisol

- establishment of normal ranges for THEG in urine

- preparation of assay tubes for home testing.

The first objective consisted of 4 important steps. Failure to achieve any one of these would necessitate curtailment of the study.

a) synthesis of THEG since it can no longer be purchased commercially at reasonable cost

b) raising of a suitable monoclonal antibody

c) preparation of a THEG-lysozyme conjugate

d) establishment of an assay for THEG using the homogeneous enzymeimmunoassay system employed in the Ovarian Monitor.

All of these research objectives (1a-b) have now been achieved as summarised below.

Summary of Results

1a Chemical and enzymatic synthesis of THEG

As reported earlier THEG has been synthesised successfully by both an enzymatic and a chemical method. The product was identified by a combination of liquid chromatography on a semi-preparative HPLC column and by mass spectroscopy. The final product was identical in all spectroscopic and chromatographic respects with an available sample of authentic material. The nmr spectrum of the sample provided absolute proof of its structure.

1b Monoclonal antibody for THEG

Availability of pure THEG material meant that immunogens for antibody production could be prepared by standard coupling techniques in our laboratory and injected into mice in the first steps of monoclonal antibody production. The immunogens (THEG-BSA conjugates) were characterised by MALDI mass spectroscopy and the  samples had a substitution level of  10-20 THEG molecules per molecule of BSA.

After injection of mice with the immunogens and applying the techniques of monoclonal antibody production a clone (8C6) has been selected, identified and cultured to provide a solution of anti-THEG antibody. As part of the screening technology an ELISA assay for THEG has been established to characterise the antibodies and this serves also as a sensitive reference laboratory assay for THEG in aqueous samples such as urine. The antibody is of class IgM and has a high specificity for THEG. Figure 1 shows a typical standard curve obtained with the ELISA for THEG using monoclonal antibody 8C6

#### Figure 1

The absorbance data are transformed logarithmically and the resulting values plotted versus the concentration of pure THEG.

The units used in the characterisation can be converted into moles/24 h which is the clinically useful unit. The mid-point value (EC50) was 0.042 (g/mL or 0.28 (mol/24 h. Thus the working range of the curve covers the expected range of values encountered in humans with an appropriate dilution of the urine sample.

###### 1b.1 Assay specificity

###### Another important parameter of the antibody is its specificity. That is the number of similar steroid and steroid glucuronide structures that cross react with the active sites of the antibody. The specificity of the assay is shown in Figure 2. Initially, a range of compounds was tested such as estriol-3-glucuronide, estriol-16-glucuronide, pregnanediol glucuronide, tetrahydrocortisone (THE), tetrahydrocortisone-3-glucuronide, and testosterone. Of these only THE showed any measurable cross reactivity which is not surprising since it is the same basic structure but minus the glucuronide ring structure.

###### Figure 2

##### The effect on assay response of a variety of steroid analogues to THEG

When the assay was tested with cortisone and tetrahydrocortisol (THF), closely related structures, only a slight cross reactivity was observed. Structures such as dehydroepiandrosterone sulphate (DHEAS) showed no significant cross reactivity. Thus the antibody is highly specific for THEG.

###### 1b.2 Sensitivity of the assay

###### As shown in Figure 1 the sensitivity of the ELISA assay is in the correct range for the measurement of THEG in urine samples without excessive dilutions of the samples. The useful working range could be set between 0.1 up to 10 µg/L and considering the daily excretion of THE of around 3 mg/day a modest dilution of urine would be anticipated.

1c  Establishment of the THEG-lysozyme assay for use with the Ovarian Monitor

The pure THEG was used to acylate lysozyme by in house methodology and a mixture of lysozyme-THEG conjugates and free lysozyme was obtained. The FPLC profile showed the usual 7 conjugate peaks and a peak of unreacted lysozyme. This mixture was partially purified by precipitation with 1 M NaCl and subsequent hydrophobic interaction column chromatography on butyl-sepharose. The results are shown in Figure 3 where the first peak consists of unreacted lysozyme and the following peaks are lysozyme-THEG conjugate mixtures. For the current work there is no need to separate out the individual conjugate families.

# Figure 3

### Hydrophobic interaction chromatography of THEG-lysozyme conjugate mixture

1d Standard curve for measurement of THEG

Fraction 33 from the column (see Figure 3 above) was used for establishment of the THEG assay. The standard curve generated with 30 (L of antibody solution is shown in the Figure 4.

Figure 4

Standard curve for measurement of THEG constructed with THEG standards prepared in buffers

#### The working range was from about 10 to 200 µmol/24 h (1.5 to 30 µg/mL)  which is an order of magnitude less sensitive than might have been expected from previous work with the Monitor fertility tests. This is most likely due to the fact the IgM antibody has 5 times the number of binding sites that a normal IgG antibody has. Since only 1 of these is required to bind to the lysozyme-THEG conjugate and inhibit it the difference between the two possibilities is in fact 10 fold. Thus, there is a buffering effect and more THEG must be added to the system to prevent the inhibition than when an IgG antibody is used.

However, a 20 minute THEG assay gave a standard curve which gave sensible results for the urine samples we have available, one of which was obtained from a stressed individual. It is also possible to construct a standard curve using a 5 minute assay as intended originally but this uses more material than does the 20 minute assay. Thus at this stage of development we have utilised the 20 minute assay protocol. However, it is clear that the goal of using the Ovarian Monitor for the THEG assay has been achieved.

1d.1 Urine matrix effects

All immunoassays are plagued by matrix effects. Simply put this means that in the presence of the urine sample the assay might behave differently. In order to test for the presence or absence of a urine matrix effect a urine sample was chosen and measured in the assay at a series of dilutions. If there is a serious urine matrix effect the amount of THEG in the sample determined by the assay result would not be a linear function of dilution. As can be seen in Figure 5 a near linear response was obtained suggesting that there is in fact little interference in the assay by urine components.

Figure 5

Dilution series showing linearity of assay response with sample dilution

A similar result was also obtained with other urine samples available in the laboratory confirming the above finding.

1d.2 Further confirmation of lack of urine matrix effects with the THEG assay

A urine sample from a 6 year old male was obtained and diluted according to the time of collection (150 mL/hr). An apparent THEG concentration in the urine was determined by measuring the lysis rate ((T20 mins) for the sample without any added THEG (the so called zero standard). From this result a series of samples could be prepared with “known” concentrations of THEG. The urine dilutions covered the range of THEG concentrations corresponding to the buffer standard curve shown in Figure 4. The results of the experiment are shown in the Ffigure 6. Also included is a similar dilution series from a woman’s urine sample. The principle involved here is that if urine matrix effects are absent the standard curve generated with the diluted urines should be parallel to the buffer standard curves shown above. Also, if the nominal concentrations of THEG determined from the assay are correct then the two sets of standard curves should overlap.

Figure 6

Data showing the lack of a urine matrix effect with the IgM antibody

This was clearly the case for the two urine specimens analysed since within experimental error the two standard curves are identical to the curve produced with the control standards prepared in buffer. This is an excellent result and means that the value for the THEG concentration given by the assay is a true measure of the actual urinary value and urine matrix effects are unimportant in this assay with this antiserum.

## 1d.3 Circadian rhythm

Cortisol levels fluctuate in a cyclical fashion through a 24 hour period. The levels are highest around 8 am and lowest from 5 pm until around midnight. The levels are also low during sleep. Thus we collected urine for a 24 hour period in 4 hourly intervals from two individuals and then analysed the samples. The 24 hour urine profile is shown in Figure 7.

Figure 7

Circadian rhythm demonstrated in the urinary THEG data collected at 3 hourly intervals

The diluted samples for both subjects clearly showed a significant fall in values (twofold) from morning to night. This is what we would expect to see. Thus, although the clinical data are preliminary we have confidence that the assay is delivering sensible numbers. A sample of the THEG data obtained showing levels is shown in Table 1.

Table 1.

###### Urine series results at different times of the day and for different individuals

The units for THEG are µmol/24 h

The values must be compared over similar time intervals because of the circadian rhythm but the levels present in timed urine specimens obtained with this assay are in the range of 10  to about 100 (mol/24 h.  These levels are quit e high and therefore relatively easy to measure. They are certainly much easier to measure than the levels of urinary free cortisol.

#### Conclusion

We have achieved the first and most important milestones set in this project under Objective 1 to the extent that we have established two assay systems for THEG one of which has the potential to be converted into a point-of-care assay. No major technical hurdles appear to remain that might prevent the remaining goals being achieved. We have commenced work on objective 2 with a local chiropractor Dr Les Whitehead. Objective 3 remains to be dome. It is also clear that the Ovarian Monitor can be used to measure THEG in a point-of-care format

#### Further work

The work remaining to be achieved can be summarised in general terms as follows

- Correlate the THEG values obtained for the Monitor assay with a urinary free cortisol assay using the same urine samples. The free cortisol assay is validated against serum free cortisol values.

- Correlate the THEG values obtained for the Monitor assay with urine against a salivary free cortisol assay using the assay situated at MacQuarie University. The salivary  cortisol assay is validated against serum free cortisol values.

- Compare THEG values in urine for chiropractic patients showing stress by elecrothermographic measurements carried out by Dr Les Whitehead and a control group which do not.

#### Commercial prospects

There is no doubt that the work completed establishes that a potentially commercial test for THEG can be developed now. No other test for THEG is available or as far as we know being developed since no one else has the pure THEG material or an antibody for THEG. Thus at this stage we would have control of the market. Equally importantly the ability to measure urinary THEG rapidly and accurately could change the way cortisol measurements are carried out and the range of application of such measurements.

Some of the commercial steps that need to be carried out include;

- Preparation of a research plan with Dr Ray Hayek to investigate the effect of chiropractic manipulation on appropriate patients

- Preparation of a dried tube point-of-care Monitor version of the THEG assay

- Establishment of a prototype lateral flow test for THEG

#### Acknowledgement

We gratefully acknowledge the financial support provided by the Australian Spinal Research Foundation.

Leonard F Blackwell

Science Haven Limited

Palmerston North

New Zealand

27 May 2008